Basic Bringup: Cryorefrigerator

# Introduction

This document is an ongoing document detailing the current status of the quick test cryorefrigerator. The intent of this document is to note current problems and attempted solutions as well as providing a history for the project.

# Goals

The goal of this project is to make a cold chamber that can quickly and cost effectively reach a few temperatures. The use for this system would be quick tests of samples prior to placing them in the main cryostats that take ages to reach temperature.

# Progress

A broad strokes design was settled on for the cold chamber. A decommissioned refrigerator tube will make the outside body of the new fridge. Then by simply having a new top plate machined we will have a sealed system.

The top plate will require a hole pattern such that the PT415 by Cryomech can bolt directly to the top plate. In addition to that a series of KF50 and potentially KF25 flanges should be machined into the plate.

Going forward the most time sensitive step is to submit the job to STS. For this reason we will mock up the top plate in a CAD software and hopefully submit the job before fleshing out the details of the assembly and experiments that are intended to be run.

For CAD see the CAD folder.

Cryomech has been emailed requesting a CAD file for the PT415. Cryomech has obliged in giving the cad files under the condition that we don’t plan on taking the PT415 out of the country (a matter of national security as far as I can tell…)

# Stress analysis

Inventor professional has a built in FEM solver for stress. Using this tool it is straightforward to get a rough idea of the deflection which the plate will see under vacuum. A pressure of 15psi was applied to the surface of the plate with the vacuum jacket marked as an immovable object. The result of the calculations is shown in the image below. The takeaway is that for a ¾” austenitic stainless steel plate a maximal deflection of 0.002” is expected.

This is of course just a simulated estimation. Some limitations include gravity not being factored in, using a generic austenitic steel grade for the simulation instead of specifying 316L stainless steel (for example). Also, the weight and the added rigidity of having the flanges populated with equipment was ignored for this test.



Talk with Jan checklist

* Bulkhead clamp dimentions (I used Kurt J Leskar , but mks has slightly different dimentions)
  + <https://www.lesker.com/newweb/flanges/flanges_technicalnotes_kf_1.cfm>
  + https://www.mksinst.com/docs/UR/isokf.pdf
* Density of KF flanges is acceptable (we could squish them closer together for an extra fillange or two if that’s worth it)
* Bolt pattern for the legs, I’ve allotted four posts each 2”x2”
* Distance to inner wall of vacuum jacket – does that leave enough for radiation shield?
* O-ring groove for the PT415
* Inventor stress analysis says that 15psi will cause a maximal deflection of 0.0023” in the centre of the mounting plate. This is fine right? (I’ll ask harmen about his program as well)